| [Insert Institution Name here] Form Template | Form No.:Equipment-1 |
| --- | --- |
|  | Revision No.: |
| Title: SAMPLE EQUIPMENT INVENTORY LOG | Approval Date: |

| Name | bench top centrifuge | refrigerator | walk in freezer | autoclave | timer |
| --- | --- | --- | --- | --- | --- |
| Unique ID |  |  |  |  |  |
| Serial # | 123456 | 789123 |  |  |  |
| Location | 137 VDL |  |  |  |  |
| Manufacturer | Corning | GE |  |  |  |
| Model # | 1512 | GE-1234 |  |  |  |
| Manual Available? Y/N | Y-manuals bookshelf | Y-manuals bookshelf |  |  |  |
| Service Contract? Y/N | Y-annual with Corning | Y – 5 year Sears |  |  |  |
| Users | Tom Jones Jane Doe | Jane Doe Mary White |  |  |  |
| Calibration Parameters (speed, temp, time, weight) | Speed Temp | Temp | Temp |  | Time |
| Operating Range | 1500-2500 rpm 10-60 min | 2 – 4 degrees C |  |  | 0.5 - 60 min |
| Calibration Tolerance (amt of error) | +/- 200 rpm +/- 1 min | +/- 2 degrees C |  |  | +/- 0.5 min |
| Frequency of Calibration | annual | daily | daily | monthly | monthly |

HOW TO USE THIS TOOL:

- Record data for each piece of equipment that contributes significantly to the study.

- Post this chart in the laboratory or area where equipment is located.

- Assign responsibility to update table as needed.

WHY DO THIS?

- The table can be used to quickly identify all equipment used in the study, equipment users, operating ranges and frequency of calibration.

- Compliance with calibration frequency can prevent errors in research findings.

- Information in the table can help troubleshoot errors or unexpected findings.

- The table can be used by auditors to quickly understand critical equipment operating parameters.

| [Insert Institution Name here] Form Template | Form No.:Equipment-2 |
| --- | --- |
|  | Revision No.: |
| Title: SAMPLE EQUIPMENT MAINTENANCE FORM | Approval Date: |

| Equipment Description (name, model): | Location: |
| --- | --- |
| Equipment Serial Number: |  |
| Standard Operating Procedure for Maintenance: | Person Responsible: |

| Routine | Non-routine* | date | Initials of person doing the work | Followed SOP? Y/N | Description of Work Performed |
| --- | --- | --- | --- | --- | --- |

*  Note:  maintenance is non-routine if the item fails or is broken and needs repair

Complete for Non-Routine Maintenance

| Nature of defect | How was defect discovered? | What action was taken to repair defect? |
| --- | --- | --- |

HOW TO USE THIS TOOL:

- Record data for each piece of equipment used in the study.

- Post this chart in the area where equipment is located.

- Assign responsibility to update table as needed.

WHY DO THIS?

- The table can be used to document routine maintenance, and to identify dates during which the piece of equipment is broken or needs repair; this may signal that research findings generated prior to repair should be re-evaluated for accuracy.

- Compliance with routine maintenance frequency can prevent errors in data generated.

- The table can be used by auditors to quickly evaluate compliance with scheduled preventive maintenance of critical research equipment.

| [Insert Institution Name here] Form Template | Form No.:Equipment-3A |
| --- | --- |
|  | Revision No.: |
| Title: SAMPLE STANDARD OPERATING PROCEDURE [SOP] FOR TIMER CALIBRATION | Approval Date: |

- Purpose

To define a procedure for calibrating timers used for time-critical procedures in the laboratory.

- Responsibility

It is the responsibility of the [insert title of supervisor] to ensure training for staff that will perform this SOP.  It is the responsibility of laboratory personnel using this procedure to read, understand, receive training for, and follow the procedure described in this SOP.

- Definitions

N/A

- Equipment Identifiers & Location

All timers in room 137.

- Safety

5.1 Training for this procedure includes review of hazards and accident prevention, personal protective equipment, and other safety requirements based on potential risks associated with methods and equipment used.

- Training

6.1 Laboratory personnel will receive training and will follow appropriate document review schedule.  Training status and SOP’s are maintained within the laboratory at [insert location].

- Maintenance

7.1 Follow recommended maintenance per manufacturer’s instructions.

- Calibration

8.1 Calibrate timers at 12 month intervals.

- Procedure

  - Log onto www.time.gov

  - Select the [your laboratory location] time zone

  - Set the time or timer channel to 5 minutes.

  - Simultaneously note the time on the website and push the start button on the timer.

  - Record the website start time on the Timer Calibration Form [Equipment-4] in HH:MM:SS

  - Record the website stop time on the Timer Calibration Form in HH:MM:SS

  - When five minutes have elapsed according to the website display, push the timer stop button.

  - Subtract the timer value from the website time value and record on the Timer Calibration Form.

  - Acceptance criteria:

    - Acceptable performance is generally +/- 30 seconds for 5 minute interval

    - If timer is not accurate under these conditions it should be removed from use.

    - If timers are found to be inaccurate, a risk-analysis should be undertaken to determine if use of the inaccurate time could have resulted in inaccurate test results, to include a review of test and control results.

- References

  - Timer calibration form

  - www.time.gov

| [Insert Institution Name here] Form Template | Form No.:Equipment-3B |
| --- | --- |
|  | Revision No.: |
| Title: SAMPLE TIMER CALIBRATION FORM | Approval Date: |

| Date mm:dd:yy | Timer ID | Channel Number | Web Start HH:MM:SS | Web Stop HH:MM:SS | Difference between elapsed Web time & Timer | Acceptable Accuracy Interval | Initials | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  | +/- 30 sec |  |  |

HOW TO USE THIS TOOL:

- Use SOP as an example of any equipment calibration SOP.

- Record data in Timer Calibration form for each timer used in the study.

- Post this form in the area where equipment is located.

- Assign responsibility to calibrate timer(s) monthly and to archive old records.

WHY DO THIS?

- Regular equipment calibration can prevent measurement errors that impact research findings.

- The form can be used by auditors to quickly evaluate compliance with scheduled research equipment calibration.

| [Insert Institution Name here] Form Template | Form No.:Equipment-4 |
| --- | --- |
|  | Revision No.: |
| Title: SAMPLE REFRIGERATOR/FREEZER TEMPERATURE FORM | Approval Date: |

Equipment and Location: _________________________________________ Year: _____________

Acceptable Refrigerator Range:  _______ to _______ Acceptable Freezer Range ______ to ______

| Month: |  |  |  |  | Month: |  |  |  |  | Month: |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Day | Refrig. Temp | Freezer Temp | Initials | Comments | Day | Refrig. Temp | Freezer Temp | Initials | Comments | Day | Refrig. Temp | Freezer Temp | Initials | Comments |
| 1 |  |  |  |  | 1 |  |  |  |  | 1 |  |  |  |  |
| 2 |  |  |  |  | 2 |  |  |  |  | 2 |  |  |  |  |
| 3 |  |  |  |  | 3 |  |  |  |  | 3 |  |  |  |  |
| 4 |  |  |  |  | 4 |  |  |  |  | 4 |  |  |  |  |
| 5 |  |  |  |  | 5 |  |  |  |  | 5 |  |  |  |  |
| 6 |  |  |  |  | 6 |  |  |  |  | 6 |  |  |  |  |
| 7 |  |  |  |  | 7 |  |  |  |  | 7 |  |  |  |  |
| 8 |  |  |  |  | 8 |  |  |  |  | 8 |  |  |  |  |
| 9 |  |  |  |  | 9 |  |  |  |  | 9 |  |  |  |  |
| 10 |  |  |  |  | 10 |  |  |  |  | 10 |  |  |  |  |
| 11 |  |  |  |  | 11 |  |  |  |  | 11 |  |  |  |  |
| 12 |  |  |  |  | 12 |  |  |  |  | 12 |  |  |  |  |
| 13 |  |  |  |  | 13 |  |  |  |  | 13 |  |  |  |  |
| 14 |  |  |  |  | 14 |  |  |  |  | 14 |  |  |  |  |
| 15 |  |  |  |  | 15 |  |  |  |  | 15 |  |  |  |  |
| 16 |  |  |  |  | 16 |  |  |  |  | 16 |  |  |  |  |
| 17 |  |  |  |  | 17 |  |  |  |  | 17 |  |  |  |  |
| 18 |  |  |  |  | 18 |  |  |  |  | 18 |  |  |  |  |
| 19 |  |  |  |  | 19 |  |  |  |  | 19 |  |  |  |  |
| 20 |  |  |  |  | 20 |  |  |  |  | 20 |  |  |  |  |
| 21 |  |  |  |  | 21 |  |  |  |  | 21 |  |  |  |  |
| 22 |  |  |  |  | 22 |  |  |  |  | 22 |  |  |  |  |
| 23 |  |  |  |  | 23 |  |  |  |  | 23 |  |  |  |  |
| 24 |  |  |  |  | 24 |  |  |  |  | 24 |  |  |  |  |
| 25 |  |  |  |  | 25 |  |  |  |  | 25 |  |  |  |  |
| 26 |  |  |  |  | 26 |  |  |  |  | 26 |  |  |  |  |
| 27 |  |  |  |  | 27 |  |  |  |  | 27 |  |  |  |  |
| 28 |  |  |  |  | 28 |  |  |  |  | 28 |  |  |  |  |
| 29 |  |  |  |  | 29 |  |  |  |  | 29 |  |  |  |  |
| 30 |  |  |  |  | 30 |  |  |  |  | 30 |  |  |  |  |
| 31 |  |  |  |  | 31 |  |  |  |  | 31 |  |  |  |  |

Comments:_________________________________________________________________________________________________________________________________________________________________________________

If the temperature limits are exceeded, then consult the manual and troubleshoot for likely causes. Document these findings within the comment section and recheck temperature in 4 hours. If recheck does not indicate temperature is within acceptance criteria, contact [insert contact person], and notify Supervisor, who will then initiate a Risk Assessment Procedure.

WHY DO THIS?

- Regular equipment monitoring can prevent measurement errors that impact research findings.

- The form can be used by auditors to quickly evaluate compliance with sample storage conditions.

| [Insert Institution Name here] Form Template | Form No.:Equipment-5 |
| --- | --- |
|  | Revision No.: |
| Title: SAMPLE BALANCE/SCALE VERIFICATION FORM | Approval Date: |

Equipment Name ________________________   Equipment ID _____________________________

Location________________________________   Acceptable Range_________________________

| Date | Calibration Weight Used | Weight Observed | Acceptable? | Initials | Comments |
| --- | --- | --- | --- | --- | --- |

- If the observed weight does not meet acceptance criteria, then consult the manual and troubleshoot for likely causes.

- Document these findings within the comment section and recheck immediately.

- If recheck indicated weight has not returned to acceptable limits, then contact service provider for help.

- Notify supervisor about situation so he/she may initiate a Risk Assessment Procedure.

HOW TO USE THIS TOOL:

- Use SOP as an example of any equipment calibration SOP.

- Record data in Balance/Scale form for each balance used in the study.

- Post this form in the area where equipment is located.

- Assign responsibility to calibrate balance(s) monthly and to archive old records.

WHY DO THIS?

- Regular equipment calibration can prevent measurement errors that impact research findings.

- The form can be used by auditors to quickly evaluate compliance with scheduled research equipment calibration.

| [Insert Institution Name here] Form Template | Form No.:Equipment-6 |
| --- | --- |
|  | Revision No.: |
| Title: SAMPLE pH METER VERIFICATION FORM | Approval Date: |

Equipment Name__________________________   Equipment Location______________________

Acceptance Criteria_________________________________________________________________

| date mm/dd/yy | pH 4.0 | buffer lot#/exp date | pH 7.0 | buffer lot#/exp date | pH 10.0 | buffer lot#/exp date | initials | comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

- If acceptance criteria are not met, consult the manufacturer’s manual and troubleshoot for likely causes.  Document these findings under “Comments” and recheck immediately.

- If recheck is not acceptable, contact the manufacturer or service provider for assistance. Put a “Rejected” sticker on equipment to remove it from service.

- Notify supervisor of the situation, and he/she will initiate a Risk Assessment Procedure.

HOW TO USE THIS TOOL:

- Use SOP as an example of any equipment calibration SOP.

- Record data in pH Meter Verification form for each meter used in the study.

- Post this form in the area where equipment is located.

- Assign responsibility to calibrate meter(s) monthly and to archive old records.

WHY DO THIS?

- Regular equipment calibration can prevent measurement errors that impact research findings.

- The form can be used by auditors to quickly evaluate compliance with scheduled research equipment calibration.

Best practices for use of laboratory pipettes

Check pipette for visual dirt/dust before using and clean with 70% ethanol if needed.

Temperature, air pressure and density of the solution can affect the performance of pipettes.

Pipettes should be stored in an upright position in order to prevent liquids from running inside the shaft of the pipette- a pipette stand designed for the pipette should be used.

Disposable pipette tips should only be used one time.

When drawing up liquid, try not to immerse the pipette tip more than 1 cm below the top of the liquid.

Use a smooth, consistent motion when depressing and releasing the plunger. Slow and careful pipetting minimizes foaming and the formation of aerosols.

Aspirate and dispense the liquid 1-3 times (pre-rinse) before drawing up the liquid for final dispensing. Do this each time you change a tip and every time you increase the volume setting.

Do not turn the pipette on its side when the tip contains liquid – this may contaminate the pipette interior.

If an air bubble forms when drawing up liquid, dispense liquid back into original container, discard the tip and use a fresh tip.

After the liquid is drawn into the tip, touch the pipette tip against the side of the container as it is being withdrawn. This will remove any liquid that may have adhered to the outside of the tip.

When dispending sample, touch the tip against the side wall of the receiving vessel to ensure that entire sample flows out of the tip.

When ejecting a tip after use, ensure that the nozzle and tip are pointed away from the face, hands and arms.

Only use the tip ejector (and not your fingers) when removing used pipette tips.

Pipettes should be inspected, cleaned and calibrated from one to four times per year depending on the use.

The main causes of pipette leak are scratched or damaged tip holders and vapor pressure from organic solvents.

The tip fit is important. If it fits too loosely, the pipettor might leak and could result in significant error in pipetting accuracy.

Leading pipette manufacturers state that seals and o-rings should last a minimum of 100,000 operations (about one year with average use).

HOW TO USE THIS TOOL:

- Provide these tips to personnel who pipette samples for the research project.

- Consider creating an SOP for pipette practices, and use SOP in employee training.

- Post this form and/or in the area where pipettes are used.

WHY DO THIS?

- Consistent equipment/pipette use can prevent measurement errors that impact research findings.